DATA VALIDATION REPORT

To:			Steve Brinkman- Stoller

From:			John Garrett

Date: 			July 28, 2006

Project/Site:		CSMRI

Project No.: 		4060-360

SDG No.:		0606093

This report presents the radiological data validation for the data obtained during the field activities for the above referenced work assignment.  The purpose of this review is to provide a technical evaluation of the radiological results that were obtained by Paragon Procedure PA SOP 713R8 for gamma emitting radionuclides by gamma spectrometry by Paragon Analytics, Inc. (Fort Collins, CO).  This report consists of 22 soil samples for the CSMRI/ 4060-360 project with 9 samples plus 1 Duplicate sample collected on June 7, 2006 and 11 remaining samples plus 1 Duplicate sample collected on June 8, 2006 and 2 duplicate samples.  The Soil samples were analyzed for gamma emitting radionuclides on July 6, 2006 and July 7, 2006 by gamma spectrometry. All analyses were conducted by Paragon Analytics, Inc.  The field sample numbers and corresponding laboratory numbers are presented below:

| Laboratory Sample Number | Client Sample Number | Matrix |
| --- | --- | --- |
| 0606093-1 | 20060607-001 | Soil |
| 0606093-2 | 20060607-002 | Soil |
| 0606093-3 | 20060607-003 | Soil |
| 0606093-4 | 20060607-004 | Soil |
| 0606093-5 | 20060607-005 | Soil |
| 0606093-5A | 20060607-006 | Soil |
| 0606093-6 | 20060607-007 | Soil |
| 0606093-7 | 20060607-008 | Soil |
| 0606093-8 | 20060608-009 | Soil |
| 0606093-9 | 20060608-010 | Soil |
| 0606093-10 | 20060608-011 | Soil |
| 0606093-11 | 20060608-012 | Soil |
| 0606093-12 | 20060608-013 | Soil |
| 0606093-13 | 20060608-014 | Soil |
| 0606093-14 | 20060608-015 | Soil |
| 0606093-15 | 20060608-016 | Soil |
| 0606093-16 | 20060608-017 | Soil |
| 0606093-17 | 20060608-018 | Soil |
| 0606093-18 | 20060608-019 | Soil |
| 0606093-19 | 20060608-020 | Soil |
| 0606093-20 | 20060608-021 | Soil |
| 0606093-20A | 20060608-022 | Soil |

Data validation was conducted in accordance with the Analytical Services Statement of Work for the following modules: the Determination of Radionuclides by Gamma Spectrometry, Module RC11-v1, October 1, 2002 and U.S. DOE Quality Systems for Analytical Services rev 1 (QSAS).

The radiological data were evaluated based on the following parameters:

- Data Completeness

- Holding Times and Preservation

- Instrument Initial Calibrations

- Instrument Performance Checks

- Preparation Blanks

- Duplicate Sample Results

- Laboratory Control Samples (LCS) Results

- Laboratory Control Samples Duplicate (LCSD) Results

- Compound Quantitation and Reporting Limits (full validation only)

Data Completeness

The data package was complete as per Paragon Procedure PA SOP 713R8 for gamma emitting radionuclides by gamma spectrometry for SDG 0606093.

Holding Times and Preservation

Analytical holding times were evaluated and all criteria were met.

Calibrations

The instruments were calibrated at the required frequency.

## Initial Calibration

All instruments were calibrated properly using NIST traceable SRM.

Instrument Performance Checks

All isotopes were within criteria.

Preparation Blanks

All of the isotopes that were analyzed had activities that were below their respective MDCs in the preparation blanks.

# Duplicate Sample Analysis

All reported Duplicate isotopes were within criteria of RER < 3 and no action was required.

Matrix Spike/Matrix Spike Duplicates

Matrix spike/matrix spike duplicates were not performed for the samples in this SDG, nor were any required.

# Laboratory Control Samples

All recoveries were within 75-125% limits.  No calculation errors or transcription errors were found.

Analyte Quantitation and Reporting Limits

Analyte quantitation was evaluated for all samples.  No calculation or transcription errors were found.  The results and reporting limits were correctly reported.

The RDL was not met for Ra-228 in samples 20060607-002, 20060607-004, 20060607-005, 20060607-005A, 20060607-006, 20060608-011, 20060608-015, and 20060608-017. Since the results were not 5 times the MDC the results are qualified as Estimated (J).

# Overall Comments

The laboratory reports a significant low bias to Lead-214 (32%) and Bismuth-214 (23%) results. All Lead-214 and Bismuth-214 results are qualified as Estimated (J).

The laboratory reports that the minimum nuclide identification criteria were not met for Al-26 in sample 20060607-003, Th-234 in sample 20060607-004, Co-57 in sample 20060607-005, Co-56 in sample 20060607-005A, Eu-152 in sample 20060607-007, Ra-228 in sample 20060607-008, Ra-228 in sample 20060608-009, Bi-212 in sample 20060607-010,  Bi-212 in sample 20060607-012,  Ra-228 in sample 20060608-012, Bi-212 in sample 20060608-013, Co-56 in sample 20060608-013 Duplicate, and Th-234 in sample 20060608-013 Duplicate. The samples are reported as tentatively identified and are qualifies as Estimated (J/UJ).

The sample densities were less than the associated calibration standard density in all samples except 20060607-001, 20060608-013 and 20060608-018 resulting in potential high bias. No action was taken.

The laboratory was unable to calculate activity and uncertainty values for Al-26 in sample 20060608-018 and is qualified as Rejected (R).

DATA QUALIFIER DEFINITIONS

For the purpose of Data Validation, the following code letters and associated definitions are provided for use by the data validator to summarize the data quality.

R	-	Reported value is “rejected.”  Resampling or reanalysis may be necessary to verify the presence or absence of the compound.

J	-	The associated numerical value is an estimated quantity because the Quality Control criteria were not met.

U J	-	The reported quantitation limit is estimated because Quality Control criteria were not met.  Element or compound was not detected.

U 	-	The material was analyzed for, but was not detected above the level of the associated value.  The associated value is either the sample quantitation limit or the sample detection limit.

NR	 -	Result was not used from a particular sample analysis.  This typically occurs when more than one result for an element is reported due to dilutions and reanalyses.