DATA VALIDATION REPORT

To:			Steve Brinkman- Stoller

From:			John Garrett

Date: 			August 24, 2006

Project/Site:		CSMRI

Project No.: 		4060-360

SDG No.:		0606204

This report presents the inorganic metals data validation for the data obtained for fourteen soil samples and one Leachate collected on June 23, 2006 and June 26, 2006 and submitted to Paragon Analytics, Inc on 06 June 27, 2006 for the above referenced work assignment. The purpose of this review is to provide a technical evaluation of the inorganic metals results that were obtained by SW-846, 3rd edition, Method 6010B and Paragon Procedure PA SOP 834R4 for trace metals by Inductively Coupled Plasma (ICP) atomic emission spectrophotometry analysis, and mercury results that were obtained by SW-846, 3rd Edition, Method 7470A and PA SOP 812R11 for mercury by Cold Vapor Atomic Absorption (CVAA) analysis for SDG 0512165 from Paragon Analytics, Inc. (Fort Collins, CO). The water samples were analyzed for dissolved ICP trace metals on June 29, 2006, and dissolved mercury by CVAA on June 30, 2006, and TCLP metals and mercury on June 30, 2006.  All analyses were conducted by Paragon Analytics, Inc.  The field sample numbers and corresponding laboratory numbers are presented below:

| Laboratory Sample Number | Client Sample Number | Matrix |
| --- | --- | --- |
| 0606204-1 | 20060623-033 | Soil |
| 0606204-2 | 20060623-036 | Soil |
| 0606204-3 | 20060623-039 | Soil |
| 0606204-4 | 20060623-043 | Soil |
| 0606204-5 | 20060623-044 | Soil |
| 0606204-6 | 20060623-046 | Soil |
| 0606204-7 | 20060623-047 | Soil |
| 0606204-8 | 20060623-048 | Soil |
| 0606204-9 | 20060623-049 | Soil |
| 0606204-10 | 20060623-050 | Soil |
| 0606204-11 | 20060623-051 | Soil |
| 0606204-12 | 20060623-052 | Soil |
| 0606204-13 | 20060623-053 | Soil |
| 0606204-14 | 20060623-054 | Soil |
| 0606204-15 | 20060623-043 | Leachate |

Data validation was conducted in accordance with the USEPA Contract Laboratory Program National Functional Guidelines for Inorganic Data Review.

The metals and mercury data were evaluated based on the following parameters:

- Data Completeness

- Holding Times and Preservation

- Initial and Continuing Calibration Verification

- Contract Required Detection Limit (CRDL)

Preparation/ Initial (ICB)/ and Continuing (CCB) Calibration Blanks

- Interference Check Sample (ICSA) Results

- Matrix Spike Results

- Duplicate Sample Results

- Laboratory Control Samples (LCS) Results

Serial Dilution Sample Results

- Compound Quantitation and Reporting Limits (full validation only)

- All criteria were met for this parameter

Data Completeness

The data package was complete except for the missing CRDL (2B) and IDL (10) QC Summary Forms. No results were qualified as a result of the missing data.

Holding Times and Preservation

Analytical holding times were evaluated and all criteria were met.

The samples were not preserved. The laboratory reports that the samples were received at 12.8 degrees Celsius. No additional action was taken.

Initial and Continuing Calibration Verification

Initial and Continuing Calibration Verification standards were analyzed at the required frequency and all were within the required 90-110% limits for ICP trace metals and within 80-120% for mercury by CVAA.  No action was necessary.

Contract Required Detection Limit (CRDL)

No CRDL or CRI standard recovery summary forms (EPA Form 2b) were included in the data package.  The reviewer obtained the %Rs from the instrument raw data.  All CRDL %Rs for ICP and CRI %Rs for mercury were within 80-120% limits.  No action was necessary.

Preparation and Initial/ Continuing Calibration Blanks

Preparation and Initial/ Continuing Calibration Blanks are evaluated to assess the level of contamination in the preparation and analytical processes.

Preparation and Initial/ Continuing Calibration Blanks were prepared and analyzed at the required frequencies.

All of the blanks that were analyzed had concentrations that were below their respective Reporting Limits (RLs).

However, if blank results were above the Instrument Detection Limits (IDLs) and below the RLs, it caused the associated sample results to be qualified for contamination as estimated and non-detected [UJ 107].  If blank results were below the negate IDL and above the negate RL, it caused the associated sample results to be qualified for negative contamination as estimated [J 107].  No sample results were qualified due to blank contamination:

# Interference Check Sample (ICSA) Results

Interference Check Samples were prepared and analyzed at the required frequencies.

No aqueous concentrations of aluminum, calcium, iron, or magnesium exceeded the ICSA values in any of the samples.  No action was necessary.

Matrix Spike/Matrix Spike Duplicate Results

Matrix spike recoveries could not be evaluated for lead in sample 0606204-7 and mercury in sample 0606204-3 due to the concentrations in the samples were greater than 4X the concentration of the matrix spike added during digestion. The laboratory control sample results were acceptable and no action was taken.

All other TCLP, ICP and Mercury MS/MSD percent recoveries were within 75-125% limits.  No action was necessary.

# Duplicate Sample Analysis

All ICP and Mercury original sample/duplicate sample and MS/MSD differences were less than 20% RPD or less than the RDL for results less than (5)(RDL).  No actions were necessary.

# Laboratory Control Samples

The laboratory analyzed laboratory control samples for all metals and mercury.  All recoveries were within 80-120% limits.  No action was necessary.

# Serial Dilution Results

All %Ds were less than 10% for all ICP metals with the exception of zinc.

Analyte Quantitation and Reporting Limits

Analyte quantitation was evaluated for all samples.  No calculation or transcription errors were found.  The results and reporting limits were correctly reported.

# Overall Comments

No CRDL or CRI standard recovery summary forms (EPA Form 2b) were included in the data package.  The reviewer obtained the results from the raw data.  No action was necessary.

DATA QUALIFIER DEFINITIONS

For the purpose of Data Validation, the following code letters and associated definitions are provided for use by the data validator to summarize the data quality.

R	-	Reported value is “rejected.”  Resampling or reanalysis may be necessary to verify the presence or absence of the compound.

J	-	The associated numerical value is an estimated quantity because the Quality Control criteria were not met.

U J	-	The reported quantitation limit is estimated because Quality Control criteria were not met.  Element or compound was not detected.

U 	-	The material was analyzed for, but was not detected above the level of the associated value.  The associated value is either the sample quantitation limit or the sample detection limit.

NR	 -	Result was not used from a particular sample analysis.  This typically occurs when more than one result for an element is reported due to dilutions and reanalyses.