New Millennium Nuclear Technologies, LLP

575 Union Blvd; #102

Lakewood, CO 80228

Tel: 303-984-5788

Fax: 303-984-0711

Email: saggarwal@nmg.org

August 31st, 2006

Field NaI Gamma Spectrometry Analysis of the Colorado School of Mines Soils

For

STOLLER

105 Technology Drive #190

Broomfield 80021

Colorado

NMNT Field Lab Gamma Spectrometry Analysis of The Colorado School of Mines Soils

# Total Radium-226 was measured by gamma spectrometry using two Sodium Iodide detectors housed in modular portable 3-inch walled lead shields in conjunction with gamma analysis software installed on a laptop.  The photons used for the determination of Ra-226 are those at the energy of 186 keV.

# A calibration check for each of the NaI gamma detectors responses were performed for Ra-226 using the NIST Mixed Multiple Nuclide Radioactive block (NIST I.D. 952-30-1) and a Ra-226 source (NIST I.D. 541-12-3) before and after every days samples to verify instrument performance by monitoring the Cs-137 or Ra-226 peak drift and the instrument’s ability to identify the calibration radionuclide of interest.  Calibration certificates are attached below.

Sample Field Counting

All 1192 soil samples were field counted for total Ra-226 gamma activity (pCi/g) by Mr. Grant Charters of NMNT. Field Ra-226 activities were calculated using the full volume of each sample, which was contained in an 8 oz plastic storage container.  The typical sample weights were between 200g and 500g.  The Ra-226 field activities (pCi/g) were used to minimize the volumes of contaminated soil removed from the site by allowing a quantitative dynamic decision making ability in near real time to happen.  The activity data obtained during the profiling process at The Colorado School of Mines was reduced and transferred to site maps as part of the daily radiological cleanup status methodology.

Site Lab NaI Ra-226 Daily Calibration Response

| Detector | Date | Cs-137 Calibration | Ra-226 Calibration | Calibration Peak | Calibration Peak | Calibration |
| --- | --- | --- | --- | --- | --- | --- |
| I.D. |  | Source | Source | Spectral | Spectral Energy | Peak Within |
|  |  | I.D. 952-30-1 | I.D. 541-12-4 | Channel # | (KeV) | Energy Range |
|  |  | 661 KeV | 186 KeV |  |  |  |
| 1 | 6/21/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 6/21/2006 |  | 150.22 - 204.15 | 1538 | 186.03 | YES |
| 1 | 6/23/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 6/23/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 6/26/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 6/26/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 6/27/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 6/27/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 6/29/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 6/29/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 6/30/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 6/30/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 7/5/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 7/5/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 7/5/2006 |  | 166.08 - 216.52 | 1047 | 186.03 | YES |
| 2 | 7/5/2006 | 587.04 - 706.44 |  | 5343 | 661.59 | YES |
| 1 | 7/6/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 7/6/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 7/7/2006 |  | 166.08 - 216.52 | 1047 | 186.03 | YES |
| 2 | 7/7/2006 | 587.04 - 706.44 |  | 5343 | 661.59 | YES |
| 1 | 7/10/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 7/10/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 7/11/2006 |  | 166.08 - 216.52 | 1047 | 186.03 | YES |
| 2 | 7/11/2006 | 587.04 - 706.44 |  | 5343 | 661.59 | YES |
| 1 | 7/12/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 7/12/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 7/14/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 7/14/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 7/17/2006 | 620.20 - 716.75 |  | 3574 | 661.61 | YES |
| 2 | 7/17/2006 |  | 150.22 - 204.15 | 1641 | 185.93 | YES |
| 1 | 7/18/2006 |  | 166.08 - 216.52 | 1047 | 186.03 | YES |
| 2 | 7/18/2006 | 587.04 - 706.44 |  | 5343 | 661.59 | YES |
| 1 | 7/19/2006 |  | 166.08 - 216.53 | 1047 | 186.03 | YES |
| 2 | 7/19/2006 | 587.04 - 706.45 |  | 5343 | 661.59 | YES |
| 1 | 7/21/2006 | 620.20 - 716.77 |  | 3574 | 661.61 | YES |
| 2 | 7/21/2006 |  | 150.22 - 204.17 | 1641 | 185.93 | YES |
| 1 | 7/24/2006 | 620.20 - 716.78 |  | 3574 | 661.61 | YES |
| 2 | 7/24/2006 |  | 150.22 - 204.18 | 1641 | 185.93 | YES |
| 1 | 7/25/2006 | 620.20 - 716.79 |  | 3574 | 661.61 | YES |
| 2 | 7/25/2006 |  | 150.22 - 204.19 | 1641 | 185.93 | YES |
| 1 | 7/26/2006 |  | 166.08 - 216.57 | 1047 | 186.03 | YES |
| 2 | 7/26/2006 | 587.04 - 706.49 |  | 5343 | 661.59 | YES |
| 1 | 7/27/2006 | 620.20 - 716.81 |  | 3574 | 661.61 | YES |
| 2 | 7/27/2006 |  | 150.22 - 204.21 | 1641 | 185.93 | YES |
| 1 | 7/28/2006 | 620.20 - 716.82 |  | 3574 | 661.61 | YES |
| 2 | 7/28/2006 |  | 150.22 - 204.22 | 1641 | 185.93 | YES |
| 1 | 7/31/2006 | 620.20 - 716.83 |  | 3574 | 661.61 | YES |
| 2 | 7/31/2006 |  | 150.22 - 204.23 | 1641 | 185.93 | YES |
| 1 | 8/1/2006 | 620.20 - 716.84 |  | 3574 | 661.61 | YES |
| 2 | 8/1/2006 |  | 150.22 - 204.24 | 1641 | 185.93 | YES |
| 1 | 8/2/2006 |  | 166.08 - 216.62 | 1047 | 186.03 | YES |
| 2 | 8/2/2006 | 587.04 - 706.54 |  | 5343 | 661.59 | YES |
| 1 | 8/3/2006 | 620.20 - 716.86 |  | 3574 | 661.61 | YES |
| 2 | 8/3/2006 |  | 150.22 - 204.26 | 1641 | 185.93 | YES |
| 1 | 8/7/2006 | 620.20 - 716.87 |  | 3574 | 661.61 | YES |
| 2 | 8/7/2006 |  | 150.22 - 204.27 | 1641 | 185.93 | YES |
| 1 | 8/8/2006 | 620.20 - 716.88 |  | 3574 | 661.61 | YES |
| 2 | 8/8/2006 |  | 150.22 - 204.28 | 1641 | 185.93 | YES |